## Slide 1: Insertion Gain

## Slide 2: Enter Insertion Gain Mode
Press [START/STOP] on remote module

## Slide 3: Enter Quik-Probe II Menu
Press [MENU] on remote module
Highlight CREATE TARGET using arrow keys
Select INS. GAIN using arrow keys

## Slide 4: Enter Audiogram Entry Screen
Press [START/STOP] on remote module
Enter audiogram using arrow keys
NAL=NR

## Slide 5: Return to Insertion Gain Screen
Press [START/STOP] on remote module
Insertion Gain Target

## Slide 6: Position Patient for Real-Ear Test
Speaker should be 12” from patient @ 00 or 450

## Slide 7: Set up microphone for Real-Ear
Reference microphone on top of ear
Probe microphone carefully inserted
Note: Leveling is performed only with reference microphone, but it is usually easier to position both microphones at the same time

## Slide 8: Level Sound Field Speaker
Press [LEVEL] on remote module

## Slide 9: Manual Mode
The following measurements are done in the COMP/Man mode. Use [>] key to change to [Manual mode].

## Slide 10: Unaided Response (REUG)
Press [START/STOP] to begin unaided measurement and again when satisfied
In Man mode, press [unaided] to save REUG
REUG

## Slide 11: Insert hearing aid into ear
Be careful not to move the probe tube!
Hearing aid
in place

## Slide 12: Aided Response @ 50 dB SPL
Press [START/STOP] to start signal
Use arrow keys to adjust source to 50 dB SPL
Press [START/STOP] to stop signal
REAG @ 50 dB

## Slide 13: Store 50 dB Response
Press [AIDED] + [STORE] to store
Select REIR and REF 1 with arrow keys
Press [START/STOP] to store
Press [MENU] to resume testing

## Slide 14: Aided Response @ 65 dB SPL
Press [START/STOP] to start signal
Use arrow keys to adjust source to 65 dB SPL
Press [START/STOP] to stop signal

## Slide 15: Store 65 dB Response
Press [AIDED] + [STORE] to store
Select REIR and REF 2 with arrow keys
Press [START/STOP] to store
Press [MENU] to resume testing

## Slide 16: Aided Response @ 80 dB SPL
Press [START/STOP] to start signal
Use arrow keys to adjust source to 80 dB SPL
Press [START/STOP] to stop signal

## Slide 17: Store 80 dB Response
Press [AIDED] + [STORE] to enter store menu
Select REIR and REF 3 with arrow keys
Press [START/STOP] to store

## Slide 18: Store REAG Target
Select TARG and REF 4 with arrow keys
Press [START/STOP] to store
Press [MENU] to resume testing

## Slide 19: Enter Quik Probe II Menu
Press [MENU]
Highlight MULTICURVE with arrow keys
Press [START/STOP]

## Slide 20: Multicurve Screen
Highlight MULTIPLE CURVES
Press [START/STOP]

## Slide 21: View Curves
Press [MENU] twice to return to Quik Probe II Menu
50
65
80
NAL=NR

## Slide 22: Switch to PURE TONE Mode for RESR90
Use arrow keys to highlight MODE
Select PURE TONE

## Slide 23: Switch to SPL
Press down arrow key to highlight GAIN(G)/SPL
Press left arrow key to select SPL

## Slide 24: Enter Operation Parameters
Highlight OPERATION PARAM
Press [START/STOP]

## Slide 25: Select 90 dB SPL Source
Use arrow keys to select 90 dB SIGNAL SOURCE
Also select OUTPUT LIMITING

## Slide 26: Return to IG Screen
Press [MENU] twice to return to testing screen

## Slide 27: Aided Saturation Response
Press [SWEEP/START] to take pure tone sweep measurement
•
•
•
•
•
[•] = LDL in SPL
RESR90

## Slide 28: DSL i/o targets in WinCHAP

## Slide 29: DSL targets & measurements