## Slide 1: Real-ear Measurements of Directional Microphones

## Slide 2: REAR Measures to Verify Directional Performance (Revit, 2000)
Reference microphone OFF.
Measuring only at 00 does not provide information on directional performance.
Need to make 2nd measure with the real-ear loudspeaker in the “null.” (e.g., Phonak AZ and Senso at 1800; Prisma and Signia @ 110 or 2500)
Hearing aids with WDRC may present a potential problem because gain increases when the signal is in the null (gain increase as input level decreases). Thus, the “back” measure may not display the ”full” potential of attenuation from the rear.

## Slide 3: Phonak 211 dAZ: Right Ear
1350
00

## Slide 4: Testing Directional Aids in the 6050 Sound Chamber

## Slide 5: Two measurements
One measurement with the aid pointed towards the sound chamber speaker
Other measurement with the aid pointed away from sound chamber speaker

## Slide 6: Notes
The following measurements are made with the new 6050 chamber that has a flat testing area that makes it easy to position the aid correctly. However, you can do the same tests with a standard 6020 chamber.
The 6500 sound chamber(s) are not anachoic, so these measurements are not exact. However, you can still get a good idea of how the directionality of the hearing aid is working
If possible, perform measurements with the sound chamber open. However, you should still get decent results even if you have to close the sound chamber because of ambient noise

## Slide 7: Sound chamber speaker
In both the old and the new style 6500 sound chambers, the speaker is mounted on the right side of the chamber at an angle so that it is pointed at the left side of the chamber.
speaker

## Slide 8: Position aid for “front” test
Position the hearing aid so that the aid is pointed at the speaker of the sound chamber. The top microphone should be positioned at the reference test point of the sound chamber

## Slide 9: This aid is digital so…
Press [MENU] and [*] to enter Star Menu
Select Digital Speech In Noise and press [START/STOP]

## Slide 10: “Front” response of aid

## Slide 11: Position aid for “back” test
Position aid so that it points away from the sound chamber speaker. The top hearing aid microphone should still be at the reference point of the sound chamber.

## Slide 12: “Back” response of aid

## Slide 13: Comparison of back & front responses
Front response
Back Response

## Slide 14: Difference of curves