## Slide 1: Proper Placement of Hearing Aids in the Sound Chamber

## Slide 2: Testing a BTE
Be sure microphone is in the
center and facing to the right
Battery Pill
to measure battery
drain
HA-2 2cc Coupler

## Slide 3: Testing an ITE
Be sure that all
“leaks” are sealed.
If not, will have
excessive gain at
~500 Hz
HA-1 Coupler

## Slide 4
HA-1 Coupler

## Slide 7
Receiver port visible in center
of the hole

## Slide 8
Need to plug the vent from
the outside!
Be sure all possible pathways
of sound are blocked!
If not, peak at ~ 500 Hz

## Slide 9
Vent or other pathways of aid not completely
sealed!
All pathways for amplified
sound sealed!

## Slide 10: Testing a CIC
Be sure that all
possible “leaks”
are sealed
CIC Coupler + CIC software

## Slide 11: CIC Option
CIC aid in a CIC coupler (0.4-cc)
Press CIC Button

## Slide 12: Multi-Curve Comparison
CIC coupler and software versus 2-cc coupler
CIC Coupler/
CIC software
CIC Coupler/
Software Not Used

## Slide 13: Probe Module
Menu
Print
Level
Aided
Unaided
Store
Clear
Monitor
Headset
Probe
Assembly
Foot
Switch
Data/
Convert
Start/
Stop
Sweep/
Start
Up/Down
Left/Right

## Slide 14: Placement of Reference and Probe Microphones
Marked 30 mm
from the Tip for Adults
and 25 mm for children
Above the Pinna
Probe Mic
Ref
Mic
Inter-Tragal
Notch