## Slide 1: HEARING INSTRUMENTS
SPECIFICATION
MEASUREMENT
Wayne J. Staab, Ph.D. Phoenix, Arizona, USA
©2001 Wayne J. Staab, Ph.D.

## Slide 2: Hearing Aid Standards and Analysis
ANSI (American National Standards Institute)
IEC (International Electrotechnical Commission)
  Standards for measurement of air-conduction HAs
    ANSI Standard S3.22–1996
    IEC 118-7 1983
  Standards include:
  1. How to measure
  2. What performance to measure
  3. Allow for a comparison of hearing aids
Test Equipment = hearing instrument analyzers

## Slide 3: Couplers (artificial ears)
Normal ear canal has about 4-cc volume
Ear with earmold or hearing aid is about 2-cc volume between earmold and TM
Types of 2-cc couplers:
  HA-1 (for ITE and ITC aids)
  HA-2 (for BTE, eyeglass, and body aids)

## Slide 4: Hearing Aid Couplers
Sound Bore
Slit Leak
Mic
2cc
Eardrum
HA-1
HA-2

## Slide 5: HA-3 and HA-4 2cc Couplers

## Slide 6: Definitions
Reference Sound Pressure Level
    = 20µPa (0.0002 dynes/cm2)
Input Sound Pressure Level
    = Level at the microphone of the HA
Acoustic Gain
    = Difference between what is put into vs. what comes out of the HA
AGC Hearing Aid
    = Automatic Gain Control
AGC Knee Point
    = Where a linear input/output curve deviates from linearity by 2 dB
Directional Hearing Aid
    = Hearing aid with 1 mic having 2 openings

## Slide 7: Basic Formula
INPUT + GAIN = OUTPUT

## Slide 8: Calculations
Input + Gain = Output
60 30 ?
60 ? 100
? 50 110
70 50 ?
80 50 ?
+
+
+
+
+
=
=
=
=
=

## Slide 9: Hearing Aid Graphs
These are obtained from a hearing aid as measured in a hearing aid analyzer
General representation
  One graph = Gain in dB
  Another graph = Output in dB SPL
But can vary considerably

## Slide 10: Basic Hearing Aid Graph Formats
OSPL90

## Slide 11: 1. Measurement of OSPL90
OSPL90 (SSPL90)
  Output (Saturation) Sound Pressure Level
  Using a 90 dB input level to the HA mic
  VC = Full-on
  Measure over range 200–5000 Hz
Forces HA to show the loudest sound it can generate
How use?
  May relate to HA matrix (1st number), i.e., 114/40/15
  Relates directly to UCL when expressed in SPL

## Slide 12: OSPL90 Graph
80
100
110
120
90
dB SPL
.25
.5
1kHz
2
4
8
OSPL90

## Slide 13: OSPL90 Description
HFA OSPL90
  High-Frequency Average OSPL90
  Calculated from previous curve (OSPL90)
Average of:
  1000 Hz
  1600 Hz
  2500 Hz

## Slide 14: 2. Measurement of Full-On Gain
Full-on Gain
  60 dB SPL input to HA mic for non AGC aids
  50 dB SPL input to HA mic for AGC aids if they overload, and RTG position
  VC = Full-on
  Measure over range 200–5000 Hz
Measure acoustic gain of the HA
How Use?
  Middle number of HA matrix (114/40/15)
    (Last number of the matrix is the rise in the curve in dB between 500 Hz and the first peak on the graph)

## Slide 15: Full-On Gain Curve
0
20
30
40
10
dB Gain
.25
.5
1kHz
2
4
8
dB Gain

## Slide 16: Full-On Gain Description
HFA Full-on Gain
  High-Frequency Average Full-on Gain
  Calculated from previous curve
Average of:
  1000 Hz
  1600 Hz
  2500 Hz

## Slide 17: SPA Gain
Special Purpose Hearing Instrument Average (SPA)
  Used for very high frequency response hearing aids having essentially no gain at 1000 or 1600 Hz
Manufacturer determines the 3 frequencies to use for these to calculate a SP Average

## Slide 18: 3. Reference Test Gain Position
RTGP
  Reference Test Gain Position
  Reduced VC setting to show better how aid performs at a level closer to speech
    Conversational speech of 65 dB SPL with ± 12 dB variation for speech sound levels = 77 dB
HFA OSPL90 – 17 dB = RTGP
  Allows for distortion-free speech

## Slide 19: Measurements made at RTGP
Reference Test Gain
  Input = HFA OSPL90 - 77
Frequency Response Curve
  Input = 60 dB SPL
  Input = 50 dB if aid would overload: Low gain aids and AGC aids
Frequency Range
  Determined from Frequency Response Curve
  Average of 1000, 1600, and 2500 Hz, and subtract 20 dB.
  Draw horizontal line. Where horizontal line crosses response curve at left is F1 (lowest frequency) and on right is F2 (highest frequency)
AGC Hearing Aids (with 50 dB input)

## Slide 20: ANSI Curve Descriptions

## Slide 21: Frequency Range Calculation
0
20
30
40
10
dB Gain
.25
.5
1kHz
2
4
8
dB Gain
28 + 33 + 37 = 98/3 = 32.67 or 33
33 - 20 = 13
33
13
F1 = 300 Hz
F2 =4500 Hz

## Slide 22: Measurements made at RTGP (cont.)
Harmonic Distortion
  Occurs when new frequencies are generated that are harmonics of the original signal
  Input = 70 dB SPL and measure total harmonic distortion at 500 and 800 Hz
  Input = 65 dB SPL and measure total harmonic distortion at 1600 Hz
Equivalent Input Noise Level = (Ln: EIN)
  Input = 60 dB SPL, determine average coupler SPL’s at 1000, 1600, and 2500 Hz
  Remove the 60 dB SPL input and measure inherent noise of the HA
Battery Current
  Input = 65 dB SPL at 1000 Hz

## Slide 23: AGC Measurement
Measurement for AGC Hearing Aids
A. Input-Output Characteristics
  Input = measure SPL in coupler from 55 to 90 dB SPL in 5 dB steps
B. Dynamic AGC Characteristics
  Input change jumping between 55 and 80 dB
    Attack Time = time where level has stabilized to within 3 dB of steady state value for 90 dB signal
    Release Time = time where level has stabilized to within 4 dB of steady state value for 55 dB signal

## Slide 24: Time Constants (AGC Hearing Aids)
3
d
B
O
u
t
p
u
t
:
A
m
p
l
i
f
i
c
a
t
i
o
n
I
n
p
u
t
:
N
o
A
m
p
l
i
f
i
c
a
t
i
o
n
A
t
t
a
c
k
T
i
m
e
R
e
c
o
v
e
r
y
T
i
m
e
4
d
B
T
i
m
e
C
o
n
s
t
a
n
t
s
S
o
f
t
S
o
u
n
d
s
L
o
u
d
S
o
u
n
d
s
S
o
f
t
S
o
u
n
d
s

## Slide 25: Gain and Output (OSPL90) Curves
80
100
110
120
90
dB SPL
.25
.5
1kHz
2
4
8
SSPL90
0
20
30
40
10
dB Gain
.25
.5
1kHz
2
4
8
dB Gain

## Slide 26: Gain and OSPL90 Curves on Same Graph
dB SPL
.25
.5
1kHz
2
4
8
100
120
130
110
90
40
60
70
50
30
dB Gain
Zero Level:
20

## Slide 27: Coupler Gain in SPL
dB SPL
.25
.5
1kHz
2
4
8
100
120
130
110
90
Zero Level:
80

## Slide 28: OSPL90 and Gain Curve Variations
10
30
40
50
20
110
100
90
80
dB
Gain
.25
.5
1kHz
2
4
8
dB
SPL
.25
.5
1kHz
2
4
8
10 dB
Slopes
15
20
Output
Gains
Matrix: 100/35,30,25/15,20

## Slide 29: Phonak Matrices
K-Amp Class D SuperC WDRC
15/95/2•6•12 15/105/2•6•12 15/105/2•6•12 20/85/2•6•12
20/100/2•6•12 20/105/2•6•12 25/105/2•6•12 25/95/2•6•12
25/105/2•6•12 25/105/2•6•12 25/105/2•6•12 30/105/2•6•12
30/105/2•6•12 30/105/2•6•12 30/105/2•6•12 35/105/2•6•12
35/110/2•6•12 35/105/2•6•12 35/105/2•6•12 40/105/2•6•12

## Slide 30: Hearing Aid Graphs With Data
.25
.5
1kHz
2
4
8
60
80
90
100
70
80
100
110
120
90
60 dB RESP CURVE
OSPL90 CURVE
.25
.5
1kHz
2
4
8
0
20
30
40
10
50
80
100
110
120
90
130
60 dB RESP CURVE
OSPL90 CURVE
RESP dBSPL
OSPL90 dB SPL
MAX OSPL90: 113.3 dB
AT: 1500 Hz
HF AVG: 110.1 dB
HF AVG FULL ON GAIN:
AT 60 dB IN 37.3 dB
RESPONSE CURVE GAIN:
33.5 dB
REFERENCE TEST GAIN:
33.1 dB
RESP LIMIT: 73.5 dB
F1 = 330 HZ F2 = 4500 Hz
THD FREQ SRC
1.0% 500 Hz 70 dB
1.8% 800 Hz 70 dB
0.9% 1600 Hz 65 dB
EQ INP NOISE: 22.6 dB
BAT (1.3V) 0.93 mA
ANSI S3.22 - 1996
LINEAR TYPE AID
F.O.G. AT 60 dB
The primary difference between the graphical
representation of these two curves is in the
left hand scale. The curve to the left shows
the 60 dB Response Curve in dB gain,
whereas the curve to the right shows the 60
dB Response Curve in dB SPL. To equate
the two, subtract the input level of 60 dB from
the RESP dB SPL curves on the right graph.
This will provide the dB gain.

## Slide 31: IEC vs. ANSI
Subject
IEC
ANSI
RTGCP
Battery current
Input, full-on gain
Distortion
Frequency range
Tolerances
AGC measurements
Directional mic.
Ambient condition
Single frequency, 15 dB
gain control setback
60 dB SPL input
50 dB SPL if
OSPL90-SPL60 is <5 dB
1 frequency
No
No
No
No
HFA (SPA), 17 dB gain
control setback
65 dB SPL input
50 dB SPL if OSPL90-SPL60
is <4 dB
3 frequencies
Yes
Yes
Yes
Yes
Small differences
RTGCP = reference test gain control position. Redrawn from Dyrlund, 1992.
Comparison of main test methods and conditions specified in ANSI S3.22-1996 to those in IEC 118-7