## Slide 1: Measuring CROS/BICROS

## Slide 3: Head Baffle Effect Two Measures
1. Ref Mic is Disabled
2. Do Not Level
3. Probe Mic “over” the “Bad Ear”
4. Store as “Unaided” Response
FIRST MEASURE:
SECOND MEASURE:
1. Probe Mic “over” the “Good” Ear
2. Store as “Aided” Response
Difference Curve (REIG) is seen
on the Screen (next Slide) and
represents the Head Baffle Effect

## Slide 4: Head-Baffle Effect

## Slide 5: How Well Does CROS Overcome the Head Baffle Effect?
1. Ref Mic enabled and over
the “bad” ear.
2. Probe Mic in the “good”
ear.
3. Level Microphones.
4. Measure REUG and store
as “Unaided.”
5. Place receiver into good ear
and transmitter over bad ear.
6. Measure and store as Aided
7. Difference Curve (REIG)
shows benefit from CROS for
sounds arriving to the bad
ear.

## Slide 6: How Well Does BICROS Overcome Head-Baffle Effect?
1. Ref Mic enabled and over
the “bad” ear.
2. Probe Mic in the “good”
ear.
3. Level Microphones.
4. Place receiver into good ear and
turned ON; transmitter over bad
ear and turned OFF.
5. Measure and store as “Unaided.”
6. Turn Transmitter ON; Store As
Aided
7. Difference Curve (REIG)
shows benefit of adding the 2nd
mic for sounds arriving to the bad
ear.

## Slide 7: Overall REIG for CROS
First Measure: REIG of Good Ear
1. Ref Mic over Bad ear
2. Probe Mic in Good ear
3. Level and Store as Unaided
Second Measure: REAG of Good Ear
1. Aid in Good Ear
2. Store As Aided
Difference Curve (REIG) shows benefit
of the CROS Aid

## Slide 8: Overall REIG for BICROS
First Measure: REIG of Better Ear
1. Ref Mic over Bad ear
2. Probe Mic in Good ear
3. Level and Store as Unaided
Second Measure: REAG of Good Ear
1. Aid in Good Ear and BOTH Mics ON
2. Store As Aided
Difference Curve (REIG) shows benefit
of the BICROS Aid

## Slide 9: Checking for REOR for Good Ear in CROS Fits
1. Ref Mic over Good Ear
2. Probe Mic in Good Ear
3. Store as UNAIDED
4. Place earmold/aid in Good Ear
5. Aid turned OFF
6. Store as AIDED
7. Difference Curve (REIG) shows
REOR caused by earmold or
hearing aid placed in the Good
ear.