## Slide 1: Calibration of 6500-CX Coupler and Real-Ear

## Slide 2: Warning Signs a.k.a. Your 6500-CX might be out of calibration if...
The analyzer doesn’t level.
Distortion measurements seem abnormally high.
Response graph is erratic.
When the microphone is alone in the test box, the RMS OUT is significantly different than the RMS SOURCE.

## Slide 3: Three microphones with the 6500-CX
M1550E (Coupler)
M200 reference
M200 probe
Each of these microphones can be adjusted for correct calibration.

## Slide 4: M1550E (Coupler) Microphone Calibration
Calibrator
Calibrator Adapter
Wrench or bolt driver
Flathead screwdriver
You will need:

## Slide 5: M1550 Cal: Step 1
Insert adapter into calibrator
M1550E Microphone

## Slide 6: M1550E Cal: Step 2
Insert microphone into mic/adapter assembly
M1550E Microphone
Calibrator
Calibrator adapter

## Slide 7: M1550E Cal: Step 3
Turn Calibrator ON
Note calibrator frequency
and output.

## Slide 8: M1550E Cal: Step 4
Press [RESET] on the 6500-CX main module.
Note: 70 source and
70 RMS Out & Flat Gain
Curve

## Slide 9: M1550E Cal: Step 5
Press [SINE/COMPOSITE] on 6500-CX main module to put unit in Pure Tone Mode.
Noise reduction set to 4

## Slide 10: M1550E Cal: Step 6
Press [NOISE REDUCTION] on the 6500-CX main module until noise reduction is OFF
Noise reduction OFF

## Slide 11: M1550E Cal: Step 7
Check that AID OUT matches source of calibrator.
Make sure FREQUENCY matches the calibrator. Adjust with arrow keys if necessary.
(Calibrator for this measurement is 110 dB SPL at 1000 Hz)
Says “108.1 dB”

## Slide 12: M1550E Cal: Step 8
To adjust calibration, loosen GAIN lock nut on back of 6500-CX.
If AID OUT is greater than 0.5 dB from the calibrator’s designated output, the M1550E microphone is out of calibration.

## Slide 13: M1550E Cal: Step 9
With flathead screwdriver, adjust GAIN pot until AID OUT is within 0.5 dB of the calibrator’s output.
Says “110.1 dB”

## Slide 14: M1550E Cal: Step 10
Tighten the lock nut of GAIN pot on the back of the 6500-CX main module.
M1550E microphone is now calibrated!

## Slide 15: Checking reference and probe microphones together
Calibration Clip
Probe tube
You will need:

## Slide 16: Ref and Probe check: Step 1
Attach calibration clip to reference microphone

## Slide 17: Ref and Probe check: Step 2
Slip probe tube through clip so that tube rests on grill of reference microphone.
Probe tube ends here

## Slide 18: Ref and Probe check: Step 3
Put 6500-CX in real-ear insertion gain mode.
(Press [PROBE] on main module or [START/STOP] on remote module to enter IG mode.)

## Slide 19: Ref and Probe check: Step 4
Hold probe/reference assembly 12 inches (30 cm) from sound field speaker.

## Slide 20: Ref and Probe check: Step 5
Press [START/STOP] on remote module to measure the response.
The response should be a flat line at 0 dB GAIN.
In calibration
Out of calibration (15 dB
of gain in this case)
70 dB IN, but 87 dB OUT

## Slide 21: To calibrate reference and probe microphones, you will need...
Sound field calibrator
Calibrator adapter
Philips screwdriver
Probe microphone adapter

## Slide 22: Reference and Probe Calibration
Enter real-ear insertion gain mode.
Press [START/STOP] on remote module or [PROBE] on main module.

## Slide 23: Reference and Probe Calibration
Press [MENU] to enter Quik-Probe II menu.
Highlight CALIBRATE PROBE using up-down arrow keys.

## Slide 24: Reference and Probe calibration
Press [START/STOP] to enter probe calibration screen.

## Slide 25: Reference Mic Cal: Step 1
Insert reference microphone into calibrator adapter mounted on calibrator
Calibrator adapter

## Slide 26: Reference Mic Cal: Step 2
Turn calibrator ON

## Slide 27: Reference Mic Cal: Step 3
Compare reference microphone input to the calibrator’s stated output. They should be within 0.5 dB of each other.
Calibrator used for this measurement has output of 110 dB SPL.

## Slide 28: Reference Mic Cal: Step 4
If necessary, use screwdriver to make adjustments to reference pot on bottom of remote module.
Reference pot
Ref mic at correct level

## Slide 29: Probe mic cal: Step 1
Probe adapter
Fun-tak
Insert probe tube into probe adapter. Tube should stick out several millimeters.
Attach Fun-Tak to affix tube in place.

## Slide 30: Probe mic cal: Step 2
Insert probe/adapter assembly into calibrator.

## Slide 31: Probe mic cal: Step 3
Turn calibrator ON

## Slide 32: Probe mic cal: Step 4
Compare probe microphone input to the calibrator’s stated output. They should be within 0.5 dB of each other.
Calibrator used for this measurement has output of 110 dB SPL.

## Slide 33
Probe Mic Cal: Step 5
If necessary, use screwdriver to make adjustments to probe pot on bottom of remote module.
Probe mic at correct level
Probe pot

## Slide 34: All three microphones are now calibrated!

## Slide 35: Calibrating the Sound Chamber
The audio source of your sound chamber is something that may never need calibrating. However, if having trouble leveling the chamber, and the microphones are in calibration, it may be an indication that the sound chamber needs to be calibrated.

## Slide 36: Step 1
Press [RESET] to enter main coupler screen.

## Slide 37: Step 2
Press [WEIGHT] twice to put analyzer in composite unweighted mode.

## Slide 38: Step 3
Unlevel the sound chamber by placing the microphone outside of the chamber and pressing [LEVEL]

## Slide 39: Step 4
Place the microphone in the sound chamber at the reference point.
Close and latch sound chamber lid.

## Slide 40: Step 5
Verify that the RMS OUT on the video screen is within 0.5 dB of the RMS SOURCE.
Here, the RMS OUT
is 6 dB Lower than
the Source.
RMS Out (73 dB)
RMS Source (79 dB)

## Slide 41: Step 6
If the source needs adjustment, loosen the AUDIO SOURCE LEVEL lock nut on the back of the 6500-CX main module.

## Slide 42: Step 7
Using a flathead screwdriver, adjust the AUDIO SOURCE LEVEL pot on the back of the 6500-CX so that RMS OUT is within 0.5 dB of RMS Source.
79 dB In and Out

## Slide 43: Step 8
Tighten the lock nut again to keep the new source level.

## Slide 44: Calibration the sound field loudspeaker
Calibration clip
Human helper or KEMAR
You will need:

## Slide 45: Step 1
Press [PROBE] on the main module or [START/STOP] on the remote module to enter the real-ear insertion gain mode. The sound field should be UNLEVELED.

## Slide 46: Step 2
Situate person or KEMAR 12” (30 cm) from sound field speaker at a 45 degree angle.

## Slide 47: Step 3
Put calibration clip on the reference microphone.

## Slide 48: Step 4
Thread probe tube attached to probe microphone through calibration clip so that the tube rests on the grill of the reference microphone.

## Slide 49: Step 5
Place both microphones above the ear nearest to the sound field speaker.

## Slide 50: Step 6
Press [START/STOP] to start measurement. Compare the RMS SOURCE to the RMS OUT. They should be within 3 dB of each other.
RMS Source;
RMS Out.
Here, there is a difference
of 8.5 dB

## Slide 51: Step 7
To adjust the sound level of the speaker, locate the LEVEL pot on the back of the main 6500-CX module.
Loosen the lock nut.

## Slide 52: Step 8
Use a flathead screwdriver to adjust the source level until RMS OUT is within 3 dB of RMS SOURCE.
Old level
New level

## Slide 53: Step 9
Tighten the lock nut again.
Your sound field speaker is now calibrated!