## Slide 1: ALDs

## Slide 2: ALD’s Some Points to Consider
Reference microphone needs to be near the microphone of the transmitter.
Need to use higher input levels because the microphone of the FM is typically closer to the FM microphone that is the microphone in a hearing aid.
    85 dB SPL for label microphone
    Even higher for a boom microphone

## Slide 3: Real Ear FM Response
Ref Mic in position where
the FM Mic will be.
1. Turn off Environmental Mic
2. Gain/SPL: SPL
3. Composite Signal
4. Level: 80-90 dB
5. LEVEL

## Slide 4: Measure REIG of FM System
1. Place FM mic in same position
as Ref mic.
2. Place probe in ear: REUG
3. Place FM rec. in ear.
4. VC to “use-gain”
5. Measure REIG
6. Compare to REIG of HA alone