## Slide 1: Electroacoustic Analysis of Hearing Aids
 ANSI-1996
 ANSI-1992
 Profiler

## Slide 2: ANSI S3.22-1996
Multiple I/O and attack & release measurements.
Improved telecoil measurement.
AGC aids tested at reduced reference test gain.

## Slide 3: Getting to ANSI ‘96 Screen
Press [ANSI] and  button to “ANSI-96”

## Slide 4: ANSI ‘96 Options
Press [Menu]

## Slide 5: ANSI ‘96 AGC Menu

## Slide 6
Battery Pills to Measure Battery Current

## Slide 7: ANSI-96
Output Curve
Frequency Response Curve
Avg and
Max Output
Avg and
Ref Test
Gain
Freq Range
Harmonic
Distortion
Eqv Input Noise
Battery Drain

## Slide 8
Example of similar gain results
when using pure-tone and composite signals
THD Measure

## Slide 9
Example of “seeing” intermodulation
distortion with a high input level (85 dB SPL)
“Irregular” frequency response

## Slide 10: ANSI ‘96 Results
SPLITS telecoil curve: should match 50 dB response curve
Multiple I/O curves
Attack and Release Time

## Slide 11
I/O Curve @ 2000 Hz for a linear hearing aid
with output limiting at 130 dB SPL

## Slide 12
I/O Curve @ 2000 Hz for a nonlinear hearing aid
with a CR of 1.7:1
Range of Input = 40 dB
Range of output = 23 dB (82 to 105)
CR = 40/23 = 1.7

## Slide 13
Telewand

## Slide 14
Telecoil Frequency
Response

## Slide 15: ANSI S3.42-1992
ANSI-92 is different from ANSI-96.
This standard uses a composite signal (“speech-like”) instead of pure tone signals.
Four measurements (50-60-70-80 dB SPL or 60-70-80-90 dB SPL) are made in 10 dB increments to measure the change in the frequency response as a function of change in input level.

## Slide 16: Entering ANSI ‘92 Screen
Menu +ANSI Buttons
Go to screen and scroll
Down to “ANSI-92” and press [Start] on Module

## Slide 17: ANSI ‘92 Results
Linear hearing aid: no change in frequency-gain
as input level increases from 50 to 80 dB SPL and
no IMD because response at 80 dB SPL has same
“morphology” as response at 50 dB SPL
Kneepoint @
80 dB SPL

## Slide 18: ANSI ‘92 Results
Nonlinear hearing aid: change in frequency-gain
as input level increases from 50 to 80 dB SPL and
no IMD because response at 80 dB SPL has same
“morphology” as response at 50 dB SPL
CR= 2.1 (40 dB change
in input (50-90); 19 dB
(82-101 dB) change
in output.) 40/19=2.1

## Slide 19: The Profiler Option
Alternative to ANSI-92 and ANSI-96.
Uses composite or digital speech signals instead of pure tone sweeps.
Very quick.
Good for establishing a baseline.
Good for identifying characteristics of unknown hearing aids.

## Slide 20: Entering the Profiler Screen
Press [Menu]
Press * Button
Press [Start]

## Slide 21: Profiler Results
Output Curve
Distortion
Frequency Responses at
50-65-80 dB using DSPA
Eqv Input Noise
Battery Drain