FONIX® FP35

Hearing Aid Analyzer

Operator’s Manual

Ver. 3.80

Ver. 3.80: 06/11/07    ©2007 Frye Electronics, Inc.
All rights reserved

Contents

Chapter 1: An Overview of the FP35

1.1  A Guide to Using this Manual	2

1.2  Layout, Controls, & Safety	2

1.2 .1  Lamps and LCD	2

1.2.2  Front panel layout	3

1.2.3  Rear Panel Layout	4

1.2.4  Sound Chamber Connectors	5

1.3  Accessories	6

1.3.1  Standard Accessories	6

1.3.2  Real-Ear Option Accessories	7

1.3.3  Optional accessories	8

1.4  Optional Features for the FP35	12

1.5  Setup	14

1.6  Safety	14

1.7  Maintenance	16

1.7.1  Servicing the Instrument	16

1.7.2  Cleaning the Instrument	17

1.7.3  Warranty	17

Chapter 2: General Operation

2.1 Operation of Keys	19

2.1.1 Using the Function Keys	19

2.1.2 Using the Local Menus	21

2.1.3 Using the Help Menus	23

2.1.4 Using the [EXIT] and [RESET] keys	23

2.2  Opening Screen	24

2.2.1  Entering the Opening Screen	24

2.2.2  Opening Screen Display	24

2.2.3 Switching Between Settings	25

2.3  The Default Settings Menu	25

2.3.1  Customizing Your Instrument Default Settings	26

2.3.2  Changing the use of [F1] in the Opening screen	26

2.3.3  Setting the date and time	27

2.3.4 Changing the external monitor colors	27

2.3.5  Explaining the Settings in the Main Default Settings Menu	28

2.3.6  Explaining the Settings in the Advanced Menu	30

2.3.7  Explaining the Settings in the Resets Default Settings Menu	32

2.4  Aid Types and Delay Times	33

2.5 Source Types & Measurements	34

2.5.1 Understanding Pure-Tone Signals	35

2.5.1.1 Delay Settings	35

2.5.1.2 Harmonic Distortion	35

2.5.1.3 Noise Reduction	36

2.5.1.4 Warble Rates	36

2.5.1.5 Adjusting the Signal Skew	37

2.5.1.6 Three-Frequency Averages	38

2.5.2 Understanding Composite Signals	38

2.5.2.1 Filter Type	39

2.5.2.2 Intermodulation Distortion	40

2.5.2.3 Noise Reduction	43

2.5.3 Measuring RMS	43

2.5.4 Measuring the RMS of the reference microphone	45

2.5.5 Using the Earphone Jack	46

2.5.6 Using Impulse Rejection	47

2.5.7 Using an External Speaker or Sound Chamber	47

2.5.8 Measuring Telecoil	49

2.6 Printing	50

2.6.1  Choosing the Printer	50

2.6.2  Adding a Label	51

2.6.3  Changing the Thermal Paper	52

2.6.4 Fixing Paper Jams	53

2.6.5  Minimizing Fading	53

2.6.6. Using an External Printer	54

2.7 Display & Data	54

2.7.1  Data Display	54

2.7.2  Curve ID Box	55

Chapter 3: Coupler Measurements

3.1  Coupler Multicurve Display	57

3.2  Leveling the Sound Chamber	58

3.2.1 Leveling without the Reference Microphone (Standard)	59

3.2.2  Leveling Using the Reference Microphone	60

3.2.3 Saving the Leveling Information	61

3.2.4 Changing the Leveling Source	61

3.3 Hearing Aid Setup	62

3.3.1  Setting up a BTE	62

3.3.2  Setting up an ITE 	63

3.3.3  Setting up a Body Aid	64

3.4  Frequency Responses	64

3.4.1  Running a Test Curve	65

3.4.2  Running a Single Frequency Measurement	67

3.4.3 Running a Three-Frequency Average	68

3.5 Distortion Measurements	69

3.5.1 Measuring Harmonic Distortion	69

3.5.2 Measuring Intermodulation Distortion	70

3.5.3 Performing an IM Distortion Sweep	71

3.5.4 Changing the Frequencies of an IM Distortion Sweep	73

3.5.5 Turning on/off Impulse Rejection	73

3.5.6  Using the Reference Microphone	74

3.6  CIC Option	74

3.7  OES Option	75

3.8  Advanced Testing	77

3.8.1 Testing Digital Hearing Aids	77

3.8.2 Testing Directional Hearing Aids	78

3.8.3 Testing with the Reference Microphone	81

3.8.4 Reversing the Microphones	82

3.8.5 Splitting the Microphones	83

3.8.6 Changing the Signal Filters	84

3.8.7 Measuring the Telecoil Response	85

3.9 Reference	87

3.9.1 Defining the Function Keys	87

3.9.2 Explaining the Menu Items	88

Chapter 4: Automated Test Sequences	93

4.1 The ANSI Sequence	93

4.1.1  Viewing the ANSI ’96 Display	94

4.1.2  Defining the Function Keys	95

4.1.3  Leveling	96

4.1.4  Switching between ANSI 96 and ANSI 03	96

4.1.5  Setting up for ANSI testing	96

4.1.6  Running an ANSI Test	98

4.1.7  Testing Digital Hearing Aids	98

4.1.8  Understanding Menu Items	99

4.2 IEC Testing	100

4.2.1  Viewing the IEC Display	100

4.2.2  Defining the Function Keys	102

4.2.3  Setting Up the Hearing Aid for Testing	102

4.2.4  Running an IEC Test	103

4.2.5  Defining the Menu Items	103

4.3 JIS Testing	104

4.3.1  Viewing the JIS Display	104

4.3.2  Defining the Function Keys	106

4.3.3  Setting Up the Hearing Aid for Testing	106

4.3.4  Running the JIS Test	107

4.3.5 Defining the Menu Items	107

4.4 ISI Testing	108

Chapter 5: Real-Ear Measurements	111

5.1  Introduction	111

5.1.1  Real-Ear Screens	111

5.1.2  Navigation	111

5.2  Real-Ear Setup	112

5.2.1  General Setup	112

5.2.2  Internal Speaker Setup	113

5.2.3  External Speaker Setup	114

5.2.4  Placing the Probe Tube	115

5.2.5  Placing the Sound Field Speaker	116

5.2.6  Leveling the Sound Field Speaker	116

5.2.7  Testing Digital Aids	118

5.2.8  Testing Body Aids	118

5.3  Audiogram Entry Screen—Creating A Target	119

5.3.1  Entering the Audiogram Screen	119

5.3.2  Audiogram Display	119

5.3.3  Entering Audiometric Information & Creating A Target	120

5.3.4  Creating an NAL-NL1 Target	120

5.3.5  Modifying a Target	122

5.3.6  Measuring the Real-Ear to Coupler Difference	122

5.3.6.1  Analyzer Setup for RECD	122

5.3.6.2  Client Setup for RECD	123

5.3.6.3  Taking the RECD Measurement	124

5.3.7  Measuring the Real-Ear to Dial Difference	124

5.3.8  Measuring the Unaided measurement	127

5.3.9  Starting a New Client	129

5.3.10  Defining the Menu Items	129

5.4  Insertion Gain Testing	131

5.4.1  Measuring the Unaided Response	133

5.4.2  Measuring the Aided Response	134

5.4.3  Testing Directional Hearing Aids	136

5.4.4 Testing Open Fit Hearing Aids 	139

5.5  The Real-Ear SPL Screen	140

5.5.1  The SPL Display	140

5.5.2  Running an SPL Test	141

5.5.3  Important Notes on SPL Testing	142

5.5.3.1  Speech-Weighting Effects on Pure tones	142

5.5.3.2  Target Formulas 	142

5.5.4  Important Notes on DSL Testing	143

5.6  Live Speech Testing	144

5.6.1  Using Speech as a Test Signal	144

5.6.2 Measuring the occlusion effect	147

5.7  Miscellaneous Real-Ear Measurements	148

5.7.1  Measuring a Single Frequency Response	148

5.7.2  Measuring Harmonic Distortion	148

5.7.3  Measuring Intermodulation Distortion	149

5.7.3.1  To perform a static IM distortion test	150

5.7.3.2  To perform an IM distortion sweep	151

5.7.3.3  To change the frequencies used in the sweep	152

5.7.4  Using the Reference Microphone	152

5.7.5 Viewing Target Data	153

5.7.6 Measuring the Telecoil Response	154

5.8  Reference	156

5.8.1  Viewing the Real-Ear Screens	156

5.8.2 Defining the Menu Items	158

5.9  CROS, and BICROS 	162

5.9.1  Measuring the Head-Baffle Effect	163

5.9.2. Measuring the Overcoming of the Head-Baffle Effect	164

5.9.2.1 CROS	164

5.9.2.2  BI-CROS	165

5.9.3  Measuring Overall Insertion Gain 	166

5.9.3.1 CROS	166

5.9.3.2 BI-CROS	166

5.9.4  Measuring Insertion Loss to the “Good” Ear (CROS)	167

Chapter 6: 2-CC Targets	169

6.1 Target Creation	170

6.2  The Coupler Target and Coupler EarSim Displays	171

6.2.1 Viewing the Coupler Target  Screen	171

6.2.2 Viewing the Coupler EarSim Screen	173

6.3 Measurements & Adjustments	174

6.3.1 Taking Basic Measurements	175

6.3.2 Viewing Target Data	175

6.3.3 Tying the Target to a Measurement Curve	176

6.4 Reference	177

Chapter 7: The Audiometer Option	181

7.1  Introduction	181

7.2  Equipment	181

7.3  Setup	181

7.3.1  Sound Field Speaker	181

7.3.2  Insert Earphones	184

7.4  Taking Audiometric Measurements	184

7.4.1  Measurements in HL	184

7.4.2  Measurements in SPL	184

7.4.3  Measurements in Gain	185

7.5  Output Limitations	185

Chapter 8: Spectrum Analysis	187

8.1  Spectrum Analysis Measurements	187

8.2  Using the Spectrum Analysis Mode	187

8.3 Technical Details	188

Appendix A: Specifications	191

Appendix B: Software Upgrades	195

Appendix C: Calibration	199

Appendix D: Fitting Formula Tables	205

Appendix E: SPL Conversions	209

Appendix F: Digital Speech Technical Details	213