## CalCert
Manufacturer's Traceable Calibration Certificate
|  |  | Associated Sales Order: |  |  |  |  | Enter SO#
|  | Calibration Date: | Enter Cal Date | [5 empty] | Calibration Due: | #VALUE!
Equipment Type |  | Hearing Aid Analyzer |  |  | Right Transducer |  |  | Left Transducer |  | Bone
Model |  | Hearlab |  |  | ER3A |  |  | ER3A
Serial Number |  | 10807004 |  |  | 29074 |  |  | 29073
Applicable Calibration Standard: ISO 389-2:1994
To whom it may concern,
Frye Electronics certifies that the above instrument was within our manufacturing calibration standards.
We hereby certify that the subject instrument is within our specification limits in accordance with our inspections.
As Found: | x | Within Tolerance |  |  |  |  | Out Of Tolerance
As Left: | x | Within Tolerance |  |  |  |  | Out Of Tolerance
Remarks:
Frye Electronics, Inc. certifies that this instrument has been compared in accordance with the above referenced
specifications using standards with accuracies traceable to the National Institute of Standards and Technology,
derived from accepted values of physical constants, derived from ratio measurements, or compared to consensus
standards. The results contained herein relate only to the item calibrated. The calibration control system and records
are in compliance with the applicable requirements of ISO 13485.
Supporting documentation relative to traceability is available for review by appointment.
The instrument was conditioned to the laboratory environment before any measurements were taken.
The following equipment listed below was used for calibration:
Equipment |  |  | Manufacturer |  |  | Serial Number and Cal Due Date
Signal Analyzer |  |  | N/A |  |  | N/A
Sound level meter |  |  | B & K |  |  | 265196 |  | 2009-06-04
Sound level Calibrator |  |  | QC-10 Quest |  |  | QIF050066 |  | 2009-07-22
Mastoid |  |  | B & K |  |  | 1895336 |  | 2009-02-02
Calibration Technician Signature
|  | Printed Name: |  | Jana Brekke
Note on F10: Frye Employee: Enter the date of calibration
Note on F46: Select analyzer model
Note on F47: Select meter model
Note on F48: Select calibrator model
Note on F49: Select analyzer model
Note on G56: Frye Employee: Select the name of the technician who calibrated the instrument
Note on J46: Select serial number
Note on J47: Select serial number
Note on J48: Select serial number
Note on J49: Select serial number
Note on L8: Frye Employee: Enter the sales order number used when calibrating this instrument
Note on L46: Enter calibration date
Note on L47: Enter calibration date
Note on L48: Enter calibration date
Note on L49: Enter calibration date
Frye Electronics, Inc.
P.O. Box 23391
Tigard, Or. USA
97281-3391
9826 SW Tigard St.
Tigard, Or. USA
97223-5243
(503)620-2722
(800)547-8209
fax: (503)639-2479
email: sales@frye.com
Web: www.frye.com

## New Hearlab
Hearlab Calibration Data Certificate
|  | S/N: | 10807004 |  | Software: |  |  |  | Hardware:
|  | Step #2 Install OS w/BartPE CD |  |  |  | x
|  | Step #3 Max out screen resolution |  |  |  | x |  | Step #14 - Ref Microphone Calibration:
|  | Step #5 Matlab Installation |  |  |  | x |  | (110dB +/-1 dB)
|  | Step #6 ACA Software Installation |  |  |  | x |  | or
|  | Step #7 Module Activation Code |  |  |  | x |  | (114dB +/-1 dB) |  | x
|  | Step #8 Hearlab Setup Information |  |  |  | x
|  | Step #9 Turn Off All Screen Savers |  |  |  | x |  | Calibrated |  | 114
|  | Step #10 Set Background |  |  |  | x
|  | Step #11 Turn Off Automatic Login |  |  |  | x
|  | Step #12 USB Port Configuration |  |  |  | x |  | Step #15 - Speaker Calibration
|  | Step #13 Create a New Client |  |  |  | x
[8 empty] | Calibration: | x
ER3A Insert Earphone Calibration
|  | Right ER3A Serial Number: |  |  | 29074
|  | Left ER3A Serial Number: |  |  | 29073
|  | Frequency Hz | RETSPL Value dB | Microphone Correction Factors | Attenuator Setting dBHL | Std + Attn. Setting dB SPL | Referenced to 110 dB SPL. As read directly on HP35665
[8 empty] | Before |  | After
[8 empty] | Left | Right | Left | Right
|  | 125 | 26 | 0 | 70 | 96 | -14 | -13 | -12.4 | -14 | -14
|  | 250 | 14 | 0 | 70 | 84 | -26 | -24 | -23.5 | -26 | -26.2
|  | 500 | 5.5 | 0 | 70 | 75.5 | -34.5 | -33 | -39.8 | -34.5 | -34.4
|  | 750 | 2 | 0 | 70 | 72 | -38 | -38 | -37.6 | -38 | -38.1
|  | 1000 | 0 | 0 | 70 | 70 | -40 | -40 | -40.1 | -40 | -40.1
|  | 1500 | 2 | 0 | 70 | 72 | -38 | -38.5 | -37.5 | -38 | -38
|  | 2000 | 3 | 0 | 70 | 73 | -37 | -37.5 | -37 | -37.1 | -37
|  | 3000 | 3.5 | 0 | 70 | 73.5 | -36.5 | -38.5 | -38.5 | -36.5 | -36.4
|  | 4000 | 5.5 | 0 | 70 | 75.5 | -34.5 | -35.5 | -35 | -34.5 | -34.5
|  | 6000 | 2 | 0 | 70 | 72 | -38 | -39.2 | -38.5 | -37.8 | -38.1
|  | 8000 | 0 | 0 | 70 | 70 | -40 | -42.8 | -41.9 | -40.2 | -40
Bone Conduction Calibration Chart
B&K 4930 S/N 1895336 Bone S/N:
Calibration Due 02Feb2010
|  | Mastoid dB | Absolute | Absolute | Voltage | RETFL | Attenutor | Output | Frequency | HearLab output Displayed on HP35665 | Before | After
|  | re 1000 Hz | Force | Voltage | relative to 1mV | Standard | Setting | Level
|  | dB re 1 mN | mN | mV | dB re 1 mV | dB re 1 mN | dBHL | dB re 1 mV | Hz | dB V | dB V | dB V
|  | -0.5 | 0.94406 | 0.12084 | -18.3558 | 67 | 20 | 68.6442 | 250 | -51.3558
|  | -0.4 | 0.95499 | 0.12224 | -18.2558 | 58 | 40 | 79.7442 | 500 | -40.2558
|  | -0.3 | 0.96605 | 0.12365 | -18.1558 | 48.5 | 40 | 70.3442 | 750 | -49.6558
|  | 0 | 1 | 0.128 | -17.8558 | 42.5 | 40 | 64.6442 | 1000 | -55.3558
|  | 0.6 | 1.0715 | 0.13715 | -17.2558 | 36.5 | 40 | 59.2442 | 1500 | -60.7558
|  | 1.1 | 1.135 | 0.14528 | -16.7558 | 31 | 40 | 54.2442 | 2000 | -65.7558
|  | 1.7 | 1.2162 | 0.15567 | -16.1558 | 30 | 40 | 53.8442 | 3000 | -66.1558
|  | 0.5 | 1.0593 | 0.13558 | -17.3558 | 35.5 | 40 | 58.1442 | 4000 | -61.8558
|  | -7 | 0.44668 | 0.057175 | -24.8558 | 40 | 40 | 55.1442 | 6000 | -64.8558
|  | -8.8 | 0.36308 | 0.046474 | -26.6558 | 40 | 40 | 53.3442 | 8000 | -66.6558
|  |  |  | Mastoid Sensitivity [mV/N] = | 128 | = | 0.128 | mV/mN
Calculation of deflection using RETFL in accordance with ISO 389-3:1994
Calibrated with customer microphone (used or new instruments)
Calibrated with factory calibrated reference microphone.
Calibrated with factory reference sound chamber.
Enter Cal Date
| Printed Name: |  | Calibration Technician Signature | [6 empty] | Calibration Date
Note on D7: Frye Employee: Serial number pulled automatically from the Cal Cert Worksheet
Note on D72: Frye Employee: Select the name of the technician who calibrated the instrument
Frye Electronics, Inc.
P.O. Box 23391
Tigard, Or. USA
97281-3391
9826 SW Tigard St.
Tigard, Or. USA
97223-5243
(503)620-2722
(800)547-8209
fax: (503)639-2479
email: sales@frye.com
Web: www.frye.com

## TDH39 cal values (hidden)
ER3A Insert Earphone Calibration
Right ER3A Serial Number:
Left ER3A Serial Number:
Freq | RETSPL | Mic | Att | Std + | Referenced to 110 dB SPL. As read directly on HP35665
Hz | Value | Corr | Set | Att Set |  | Before |  | After
| dB | Fact | dBHL | dBSPL |  | Left | Right | Left | Right
125 | 26 | 0 | 70 | 96 | -14
250 | 14 | 0 | 70 | 84 | -26
500 | 5.5 | 0 | 70 | 75.5 | -34.5
750 | 2 | 0 | 70 | 72 | -38
1000 | 0 | 0 | 70 | 70 | -40
1500 | 2 | 0 | 70 | 72 | -38
2000 | 3 | 0 | 70 | 73 | -37
3000 | 3.5 | 0 | 70 | 73.5 | -36.5
4000 | 5.5 | 0 | 70 | 75.5 | -34.5
6000 | 2 | 0 | 70 | 72 | -38
8000 | 0 | 0 | 70 | 70 | -40

## Bone Cal Values (hidden)
Bone Conduction Calibration Chart
B&K 4930 S/N 1895336 Bone S/N:
Calibration Due 02Feb2010
Mastoid dB | Absolute | Absolute | Voltage | RETFL | Attenutor | Output | Frequency | HearLab output Displayed on HP35665 | Before | After
re 1000 Hz | Force | Voltage | relative to 1mV | Standard | Setting | Level
dB re 1 mN | mN | mV | dB re 1 mV | dB re 1 mN | dBHL | dB re 1 mV | Hz | dB V | dB V | dB V
-0.5 | 0.94406 | 0.12084 | -18.3558 | 67 | 20 | 68.6442 | 250 | -51.3558
-0.4 | 0.95499 | 0.12224 | -18.2558 | 58 | 40 | 79.7442 | 500 | -40.2558
-0.3 | 0.96605 | 0.12365 | -18.1558 | 48.5 | 40 | 70.3442 | 750 | -49.6558
0 | 1 | 0.128 | -17.8558 | 42.5 | 40 | 64.6442 | 1000 | -55.3558
0.6 | 1.0715 | 0.13715 | -17.2558 | 36.5 | 40 | 59.2442 | 1500 | -60.7558
1.1 | 1.135 | 0.14528 | -16.7558 | 31 | 40 | 54.2442 | 2000 | -65.7558
1.7 | 1.2162 | 0.15567 | -16.1558 | 30 | 40 | 53.8442 | 3000 | -66.1558
0.5 | 1.0593 | 0.13558 | -17.3558 | 35.5 | 40 | 58.1442 | 4000 | -61.8558
-7 | 0.44668 | 0.057175 | -24.8558 | 40 | 40 | 55.1442 | 6000 | -64.8558
-8.8 | 0.36308 | 0.046474 | -26.6558 | 40 | 40 | 53.3442 | 8000 | -66.6558
|  | Mastoid Sensitivity [mV/N] = | 128 | = | 0.128 | mV/mN
Calculation of deflection using RETFL in accordance with ISO 389-3:1994