7000 Measurement Methods

2004/06/11 -med & tmr

Measurement methods:

Signal   Source Frequency   Method       Window        Mode

Composite  Off   0Hz         FFT (100Hz)  Raised Cosine RE+CM

Composite  On    0Hz         FFT (100Hz)  Rectangular   RE+CM

Tone     Off   Sweep       DFT (100Hz)  Raised Cosine RE+CM

Tone     Off   100Hz **    TRMS         Rectangular   RE+CM

Tone     Off   non-100Hz*  DFT          Raised Cosine RE+CM

Tone    <85dB  Sweep       DFT (100Hz)  Rectangular   RE

Tone    <85dB  100Hz **    DFT          Rectangular   RE

Tone    <85dB  non-100Hz*  DFT          Rectangular   RE

Tone   >=85dB  Sweep       TRMS (100Hz) Rectangular   RE

Tone   >=85dB  100Hz **    TRMS         Rectangular   RE

Tone   >=85dB  non-100Hz*  DFT          Rectangular   RE

Tone    <90dB  Sweep       DFT (100Hz)  Rectangular   CM

Tone    <90dB  100Hz **    DFT          Rectangular   CM

Tone    <90dB  non-100Hz*  DFT          Rectangular   CM

Tone   >=90dB  Sweep       TRMS (100Hz) Rectangular   CM

Tone   >=90dB  100Hz **    TRMS         Rectangular   CM

Tone   >=90dB  non-100Hz*  DFT          Rectangular   CM

Coupler

Mic Cal  Off   1000Hz      TRMS         None          CM

Notes:

* non-100Hz includes all frequencies above 8000Hz and all

frequencies that are not a harmonic of 100Hz.

** 100Hz means 100Hz harmonics from 100Hz to 8000Hz.

RE = Real Ear test mode.

CM = Coupler (chamber) test mode.

FFT = Standard FFT using 256 samples over 10mS, 100Hz harmonics.

DFT = Standard DFT using variable samples and sample rate

depending on frequency. 256 samples over 10mS for 100Hz harmonic

based signals.

ERMS = Estimated RMS is done by performing an FFT, correcting for

the data calibration and filter errors, then summing the results

to obtain the RMS. This gives good results when a full spectrum

synchronous source is used, but may not be as accurate in the

presence of significant signal noise.

CRMS = Corrected RMS is done by performing a standard RMS, then

correcting the result for the selected frequency. This gives good

results when a strong pure tone signal is present, but may not be

as accurate in the presence of strong signal noise.

TRMS = True RMS is done by performing a standard RMS without

correcting for the frequency. This works well in the presence of

strong signal noise, but does not correct for microphone

calibration errors at each frequency. The calibration for TRMS is

tied to the 1000Hz calibration reference.

.pa

Microphone Calibration:

When calibrating the coupler microphone, only the microphone

sensitivity at 1000Hz is changed. The coupler microphone calibration uses the True RMS measurement method. This is done because the Standard calibration method for the coupler microphone uses 1000Hz as the 0dB sensitivity reference level, so this will provide the most accurate calibration reference for the

microphone. Using True RMS for the measurement allows frequencies

other than 1000Hz to be used as the calibration source. However,

for best accuracy, it is recommended that a 1000Hz reference

source be used for the Coupler microphone calibration.

<eof>