Special RS232 commands on the 7000

2004/06/11 -med & tmr

The 7000 can do things via RS232 that can't be done from the

keyboard. The following list describes some of the more useful

commands available.

Set Level:

The source level can be set in 0.01dB increments from 0.01dB

to the maximum possible level that the hardware can produce.

The maximum level is dependent upon the configuration and

leveling. The absolute maximum possible source Spl setting

allowed is 120.00dB. Setting the source to 0dB turns it off.

Set Frequency:

The frequency can be set in 50Hz increments from 50Hz to 8KHz.

However, it should be noted that the microphones are only

calibrated from 100Hz to 8KHz. Also it should be noted that the

7000 is designed for optimal operation between 200Hz and 8KHz.

The frequencies below 200Hz are rolled off by the hardware. Leveling and compensation is centered around the 200Hz to 8KHz range. Signals outside that range are provided with whatever level they have after the correction has been applied to the signal data. In addition, signals below 200Hz may not be as accurate. They are provided as is. The lowest possible frequency is 50Hz.

Noise Reduction:

Noise reduction can be set from 0 to 127. Care should be used

when setting large noise reduction values in pure tone because it

will cause the keyboard to become sluggish.

When in synchronous measurement (source on), Each doubling of the

noise reduction value adds an additional one bit of effective

signal resolution (6dB for each bit added).

2=6dB, 4=12db, 8=18db, 16=24db, 32=30db, 64=36db, 127=42db

Note: adding noise reduction does not increase the signal

amplitude, it improves the ability of the 7000 to measure low

level signals by reducing non-synchronous noise and providing

additional effective bits to the A2D conversion through time

averaging. Thus setting the noise reduction to 16 makes the 7000

measurements seem like it has a 16bit A2D in software even though

the hardware is only 12bits. This can help to make signals with a

large dynamic range appear smoother on the portions that have

less amplitude (near the bottoms of curves).

.pa

Impulse Rejection:

Normally, except during leveling, Impulse rejection is set to

off. You can apply impulse rejection by setting the level of

rejection you wish to apply to the measurement. The maximum

rejection allowed is 24dB. Care should be used when applying

small signal rejection values since it can make the measurements

take much longer in a noisy environment. Normal signal rejection

should be in the range between 3dB and 12dB.

Noise rejection is a bit complex in its operation. It is tied to

the prescalers and the noise reduction setting. Noise rejection

can only occur when noise reduction is on (set to 2 or higher).

After a prescale change occurs, the rejection counter is set to

zero. The next measurement made then becomes the reference level

for the signal rejection. If the rms level next measurement made

is equal to or greater than the previous rms level plus the noise

rejection level set, the measurement will be discarded. The

number of sequentially rejected measurements allowed is

determined by the Noise Rejection setting divided by two. Thus if

noise rejection is set to 4, the Impulse rejection limit is set

to 2. Meaning that after two successive measurement rejections,

the last measured signal rms will become the new impulse

rejection reference level. Immediately after an input prescale

change, the next measurement will become the new impulse

rejection reference level and the impulse rejection count is set

back to zero. Each valid (non-rejected) signal becomes the new

reference level and clears the impulse rejection count.

Impulse Rejection uses the overall rms of the last valid measured

signal. Impulse rejection works on the assumption that an

impulse will be a rms signal level that is above the previously

measured signal. A negative rms value below the reference level

is treated as a valid signal and will become the new impulse

reference level. If the input prescale (input gain) is changed,

the next measurement made will become the new impulse rejection

reference level.

.pa

Raw Signal:

Normally the measured data from the 7000 is processed and

presented as dB values. Via RS232 it is possible to get the last

raw (time domain) measurement made. This is the output from the

A2D converter. You can either get the last measurement made, or

request a new measurement to be made. You can ask for the

measurement information from either channel. If a specific

measurement is made, you can set the number of samples to be

collected in the measurement (1024 max, 16 min).

You can ask for either the raw measured data directly from the

A2D converter, or you can ask for the time averaged data

(determined by the noise reduction and impulse rejection

settings).

Signal Information:

You can also ask for the specifics of the measurement conditions

of the last measurement made.

SampleRate       = samplerate of signal

CaptureSize      = number of capture samples

SourceSamples    = number of samples in source

RampSamples      = source ramp up/down samples

CouplerSkew      = source to coupler mic sample delay

ProbeSkew        = source to probe mic sample delay

InA.PrescaleGain = InputA (coupler mic) prescale gain

InB.PrescaleGain = InputB (probe mic) prescale gain

OutL.RmsOffset   = Lchan (speaker) single tone to rms offset

OutR.RmsOffset   = Rchan (Tcoil) single tone to rms offset

OutL.MaxLevel    = Lchan (speaker) max possible output level

OutR.MaxLevel    = Rchan (Tcoil) max possible output level

InA.GainIndex    = InputA (coupler mic) prescale gain index *

InB.GainIndex    = InputB (probe mic) prescale gain index *

* GainIndex was added to the list with V2.12 software

Screen Capture:

The 7000 allows you to retrieve the current screen image. This is

useful for publication of information presented on an 7000 screen. A program (FCAPTURE.EXE) is provided with the RS232

Technical Support Information which allows you to capture a 7000

screen image to a Windows BMP file.

<eof>