µ §

Proposal on 
DataFmtCodeStd 200
for REMHIT definition

Draft 1 : 	1996-07-31 Peter Kossek, REXTON DANPLEX A/S

Draft 2 : 	1996-08-01 Peter Kossek, REXTON DANPLEX A/S

Final version:	1996-08-16 Peter Kossek, REXTON DANPLEX A/S
		no changes, accepted in Draft 2 state.

As the current REM and HIT definition (format code 100) is used differently by the AEM's, the following description serves as a proposal on exactly how to interpret data in REMHIT, with the goal of defining a format code 200 following this proposal and end up with a true public data definition handled and interpreted the same way by everybody involved in this system. There are no changes in the data format itself.

The following is based on REMHIT.PAS from noah 2.0. Extra descriptions added where found applicable. This format was chosen for easy adaptation to programming using this definition.

Draft 2 is based on draft 1 and the discussion/decisions from the meeting at HIMSA I/S on August 1, 1996.

Changes to the REMHIT format:

·	the REIR curve may now be used as both compensated and un-compensated (for REUR) identified by the value of Measurement

·	the REIR curve may also be used as a gain curve by setting the input part to zero and storing the gain in the output part

·	meas_InsertGain renamed to meas_InsertGain_C (value 5) for identification of Insertion Gain measurement compensated for the REUR curve

·	meas_InsertGain_U (value 50) added for identification of Insertion Gain measurement uncompensated for the REUR curve

To make the REMHIT format match the OnLine Protocol No: 1, following changes to the OnLine Protocol No: 1 were agreed upon:

·	itc_REIR_L renamed to itc_REIR_L_C (value 5), 
Meaning is extended to "Real Ear Insertion Response, Left Ear, with REUR compensation"

·	itc_REIR_R renamed to itc_REIR_R_C (value 6), 
Meaning is extended to "Real Ear Insertion Response, Right Ear, with REUR compensation"

·	itc_REIR_L_U (value 9) added,
 Meaning is extended to "Real Ear Insertion Response, Left Ear, without REUR compensation"

·	itc_REIR_R_U (value 10) added, 
Meaning is extended to "Real Ear Insertion Response, Right Ear, without REUR compensation"

Changes to the OnLine Protocol No: 1 document will be made by HIMSA I/S.

August 1, 1996

Peter Kossek,

REXTON DANPLEX A/S, DENMARK

Unit remhit;

(*************************************************************************

Based on REMHIT.PAS for noah version 2.0.

Updated file for public format 200.

Original data structure unchanged, interpretation and handling decriptions added.

NOTE! UndefInt can be used in any data field to indicate that a value is not applicable.

Peter Kossek, REXTON DANPLEX A/S, DENMARK

August 1,  1996.

Please remember to include the remarks when coverting this into a .H file.

**************************************************************************)

interface

uses WinTypes,WinProcs,noahdef;

{ TSignalType : The signal type used during the measurement }

Type  TSignalType=Integer;

Const st_Tone=1;

st_Warble=2;

st_NarrNoise=3;

st_TwoTone=4;

st_WhiteNoise=5;

st_PinkNoise=6;

st_SpeechNoise=7;

st_Patient=8;

st_User1=9;

st_User2=10;

st_User3=11;

{ TSignalOutput : The channel/media used to present the signal/stimulus }

Type  TSignalOutput=Integer;

Const so_InternalBox=1;

so_ExternalBox=2;

so_FF=3;

so_InternalBoxCoil=4;

so_ExternalBoxCoil=5;

so_FFCoil=6;

so_AC=7;

so_User1=8;

so_User2=9;

so_User3=10;

{ TCouplerType : The device in which the probe microphone is placed, in other words: this is where the output is read }

Type  TCouplerType=Integer;

Const ct_None=1;

ct_RealEar=2;

ct_711=3;

ct_2cc=4;

ct_FreibKK=5;		{ Freiburger Konischer Kuppler }

ct_FreibKKK=6;		{ Freiburger Konischer Kinder Kuppler }

ct_User1=7;

ct_User2=8;

ct_User3=9;

{ TBattType : The battery type used }

Type  TBattType=Integer;

Const bt_None=1;

bt_Mercury=2;

bt_ZincAir=3;

bt_OtherType=4;

bt_User1=5;

bt_User2=6;

bt_User3=7;

Type  TV1000=WORD;

TOhm1000=WORD;

TBattVoltage=TV1000;	{ The voltage of the battery in milliVolts }

TBattImp=TOhm1000;	{ The impedance of the battery in milliOhms }

TBattPill=Integer;		{ TBattPill: the type of the battery pill }

Const bp_None=1;

bp_Bat312=2;

bp_Bat13=3;

bp_Bat230=4;

bp_Bat675=5;

bp_User1=6;

bp_User2=7;

bp_User3=8;

{ TMeasMode : Which measurement mode was used during the measurement }

Type  TMeasMode=Integer;

Const mm_Sweep=1;

mm_FFT=2;

mm_TimeMeas=3;

mm_Battery=4;

mm_User1=5;

mm_User2=6;

mm_User3=7;

{ TMeasurement: Identification of the measurement, mainly for use with data export and storage of not yet public defined measurements at the end of the public buffer }

Type  TMeasurement=Integer;

Const meas_AudHand=1;			{ Audiometry }

meas_TargHand=2;			{ Target Curve }

meas_UnAided=3;			{ Unaided Response }

meas_OcclResp=4;			{ Occluded Response }

meas_InsertGain_C=5;		{ Insertion Response (compensated for REUR) }

meas_AidedResp=6;			{ Aided Response }

meas_InputOutput=7;			{ Input Output }

meas_HarmDist=8;			{ Harmonic Distortion }

meas_OcclEff=9;			{ Occlusion Effect }

meas_RECD=10;			{ Real Ear to Coupler Difference }

meas_SPL90=11;			{ XSPL-90, i.e. OSPL-90 for IEC and SSPL-90 for ANSI }

meas_FOG=12;			{ Full On Gain }

meas_FreqResp=13;			{ Frequency Response}

meas_BattCurr=14;			{ Battery Current }

meas_InterDist=15;			{ Intermodulation Distortion }

meas_EquivNoise=16;		{ Equivalent Input Noise}

meas_AttackRec=17;			{ Attack/Recovery}

meas_IndCoil=18;			{ Induction Coil }

meas_User1=19;			{ User specific #1 }

meas_User2=20;			{ User specific #2 }

meas_User3=21;			{ User specific #3 }

meas_InsertGain_U=50		{ Insertion Response (uncompensated for REUR) }

Type  TRHMeasCond= record

ManufCode:TManufCode;

DevTypeCode:TDevTypeCode;

SignalType:TSignalType;

SignalOutput:TSignalOutput;

SignalLevel:TdB10;

SignalFreq:THertz;			{ Set this to UndefInt for multifrequency measurements,

otherwise state the frequency here}

BattType:TBattType;

BattPill:TBattPill;

BattVoltage:TBattVoltage;

BattImp:TBattImp;

UseRECoupler:BOOL; 		{ TRUE, if measurement done using Real Ear to Coupler Difference }

MeasMode:TMeasMode;

Measurement:TMeasurement;

end;

TFreqMeasPoint= record

Freq:THertz;			{ The frequency at which Input/Output was recorded }

Input:TdB10;			{ Input value }

Output:TdB10;			{ Output value }

end;

{ TFreqMeas: generic two channel frequency response type measurement }

TFreqMeas= record

MeasCond:TRHMeasCond;

FreqCrv:Array[0..169-1] of TFreqMeasPoint;

end;

{ TOcclMeas: Occlusion Effects, comparison between the response of the un-occluded ear with the reponse of the 
	occluded ear .

OpenEarCrv is the response of the un-occluded ear, no compensation in input or output. 
	OccEarCrv is the response of the occluded ear, no compensation in input or output .}

TOcclMeas= record

MeasCond:TRHMeasCond;

OpenEarCrv:Array[0..169-1] of TFreqMeasPoint;

OcclEarCrv:Array[0..169-1] of TFreqMeasPoint;

end;

{ TRECMeas : Real Ear to Coupler Difference measurement

RECDCrv is to be read as gain (output - input), where the result gives the difference between a measurement 	in a coupler and a measurement in the client's ear. As both parts of this measurement are performed with the 	same measurement conditions, the curve can be used in a manner with the input part holding the coupler SPL 	output curve, and the output part holding the real ear SPL output curve. This will yield the correct gain 	value.}

TRECMeas= record

MeasCond:TRHMeasCond;

RECDCrv:Array[0..169-1] of TFreqMeasPoint;

end;

{ TFittingRule : The fitting rule used for the calculation of a target curve }

TFittingRule=Integer;

Const fr_POGO=1;

fr_POGOII=2;

fr_NAL=3;

fr_NALProf=4;

fr_Berger=5;

fr_HalfGain=6;

fr_ThirdGain=7;

fr_DSL=8;

fr_LIBBY=9;

fr_Byrne=10;

ft_CoxMSU=11;

fr_User1=100;

fr_User2=101;

fr_User3=102;

fr_User4=103;

fr_User5=104;

fr_User6=105;

fr_User7=106;

fr_User8=107;

fr_User9=108;

fr_User10=109;

fr_Undefined=UndefInt;

{ THIType : The type of the hearing instrument being tested }

Type  THIType=Integer;

Const hit_ITE=1;

hit_BTE=2;

hit_ITC=3;

hit_MITC=4;

hit_Body=5;

hit_User1=6;

hit_User2=7;

hit_User3=8;

hit_User4=9;

hit_User5=10;

hit_Undefined=UndefInt;

Type  Tmm10=Integer;

{  TTargetPoint : a single element/point of the target curve }

TTargetPoint= record

TargetFreq:THertz;

TargetGain:TdB10;

end;

{ TTargetCurve : a full target curve including the description }

TTargetCurve= record

ManufCode:TManufCode;

DevTypeCode:TDevTypeCode;

FittingRule:TFittingRule;

HIType:THIType;

VentDiam:Tmm10;				{ The diameter of the vent canal }

VentLen:Tmm10;				{ The length of the vent canal }

ResGain:TdB10;				{ The reserve gain included in the target curve }

CouplerType:TCouplerType;

SignalLevel:TdB10;				{ Use this field to state the signal level if target curve relates to a 						   specific signal level }

Target:Array[0..24-1] of TTargetPoint;	{ The target curve }

RuleName:Array[0..51-1] of char;		{ Text containing the target curve's name/description }

end;

TPct100=Integer;

{ TTHDDistMeasPoint: a single measurement point for Harmonic Distortion Measurements. }

TTHDDistMeasPoint= record

Freq:THertz;			{ Signal frequency }

Input:TdB10;			{ Signal level }

Output1Harm:TdB10;		{ the output value at Freq }

Output2Harm:TdB10;		{ the output value at 2*Freq }

Output3Harm:TdB10;		{ the output value at 3*Freq }

THDPct:TPct100;			{ the calculated THD based on OutputXHarm }

end;

{ TTHDDistMeas: a complete curve containing a Harmonic Distortion measurement.}

TTHDDistMeas= record

MeasCond:TRHMeasCond;

DistCrv:Array[0..161-1] of TTHDDistMeasPoint;

end;

{ TTIMDistMeasPoint: a single measurement point for InterModulation  Distortion Measurements.}

TTIMDistMeasPoint= record

Freq1:THertz; 		{ First stim freq }

Freq2:THertz;		{ Second stim freq }

Input1:TdB10;		{ Level of first stim }

Input2:TdB10;		{ Level of second stim }

Output1:TdB10;		{ Output at Freq1 }

Output2:TdB10;		{ Output at Freq2 }

OutputDif1:TdB10;		{ Output at Freq2-Freq1 }

OutputDif2:TdB10;		{ Output at 2*Freq1-Freq2 }

TIMPct:TPct100;		{ The calculated result }

end;

{ TTIMDistMeas: a complete curve containing a Intermodulation Distortion measurement.}

TTIMDistMeas= record

MeasCond:TRHMeasCond;

DistCrv:Array[0..161-1] of TTIMDistMeasPoint;

end;

{ TIOMeasPoint : a measurement point for Input/Output measurements. }

Note: no frequency information in single points. This is common for all the measurement points and the value is stated in MeasCond.}

TIOMeasPoint= record

Input:TdB10;

Output:TdB10;

end;

{ TIOMeas : a complete curve containing an Input/Output measurement }

TIOMeas= record

MeasCond:TRHMeasCond;

IOCrv:Array[0..61-1] of TIOMeasPoint;

end;

TmA100=WORD; 		{ Current value in hundreths of a milliampère }

{ TBatMeasPoint : a measurement point for Battery Current measurements. }

TBatMeasPoint= record

Freq:THertz;

Current:TmA100;

end;

{ TBatMeas : a complete curve containing a Battery Current measurement. }

TBatMeas= record

MeasCond:TRHMeasCond;

BatCrv:Array[0..169-1] of TBatMeasPoint;

end;

TARTime=Integer; 	{ Attack/Release time in milliseconds }

{ TARMeasPoint : a measurement point for Attack/Release measurements }

TARMeasPoint= record

Output:TdB10;

end;

{ TARCurve : a single curve for partial info of an Attack/Release measurement }

TARCurve= record

Curve:Array[0..256-1] of TARMeasPoint; {The curve itself }

Result:Integer;	{ Time result (in milliseconds) (result depending on if this is AttackCrv or ReleaseCrv)}

Resolution:Integer;	{ The time resolution of the measurement (in milliseconds) }

Predelay:Integer;	{ The number of milliseconds included in the curve before level change }

end;

{ TARMeas : a complete curve set containing a full Attack/Release test }

TARMeas= record

MeasCond:TRHMeasCond;	{ Generic measurement conditions }

LevelStep:TdB10;		{ The level step size from the start level defined in MeasCond }

AttackCrv:TARCurve;	{ The curve where the stimulus level increases }

ReleaseCrv:TARCurve;	{ The curve where the stimulus level decreases }

end;

EINMeas= record

MeasCond:TRHMeasCond;

EINCrv:Array[0..169-1] of TFreqMeasPoint;	{ Input part: uncompensated level in the from the test chamber }						{ Output part: coupler output level compensated with the gain 						    	   value }

EINRMS:TdB10; 				{ The calculated result }

end;

RemData= record

Target:Array[0..3-1] of TTargetCurve;

{ Unaided Reponse: input is the input level, output is uncompensated probe level. Must be valid for use as an   
        	output curve as well as a gain curve, so if only output has been measured, fill the input part with the stimulus                      	level value }

REUR:TFreqMeas;

{ Occluded Response: input is the input level, output is uncompensated probe level. Actually like Aided Response 	(REAR) but the H.I. is off or detached during the measurement }

REOR:TFreqMeas;

{ Insertion Response:	input is the input level, output is the probe level with or without REUR compensation.

Measurement = meas_InsertGain_U -> the output values are dB SPL values without any REUR 	compensation.

Measurement = meas_InsertGain_C -> the output values are dB SPL values less the REUR's gain , i.e.with	REUR compensation.

Gain curves may be stored here. In this case the Input part must be set to zero and the gain value must be 	stored in the Output part. }

REIR:Array[0..5-1] of TFreqMeas;

{ Aided Response: input is the input level, output is uncompensated probe level. }

REAR:Array[0..5-1] of TFreqMeas;

IOMeas:Array[0..5-1] of TIOMeas;

HarmDistortion:Array[0..3-1] of TTHDDistMeas;

{ See the TOcclMeas description }

Occlusion:Array[0..3-1] of TOcclMeas;

{ See the TRECMeas description }

RECoupler:TRECMeas;

end;

HitData= record

SPL90:Array[0..2-1] of TFreqMeas;

FullOnGain:Array[0..2-1] of TFreqMeas;

FreqResp:Array[0..2-1] of TFreqMeas;

Battery:Array[0..2-1] of TBatMeas;

HarmDistortion:Array[0..2-1] of TTHDDistMeas;

InterDistortion:Array[0..2-1] of TTIMDistMeas;

EquivInputNoise:Array[0..2-1] of EINMeas;

IOMeas:Array[0..2-1] of TIOMeas;

AttackRecover:Array[0..4-1] of TARMeas;

InductionCoil:Array[0..2-1] of TFreqMeas;

end;

implementation

end.